Leveling the 6m antennas

Dick Plambeck, revision A, 6 Feb 2006

Tiltmeters

A Schaevitz model LSOC-1º inclinometer is installed on the azimuth platform of each 6m antenna usually near the stowpin.   Usually the inclinometer is covered with a sheet metal box.

The Bendix connector on the inclinometer is wired to the tilt1 input of the a/d mux board in the cabin, as given in the table below.  Raising the right side of the tiltmeter increases the a/d reading.  The tilt is given in arcminutes on the BIMA drives rtd screen.

| Bendix | D9 | signal |
| --- | --- | --- |
| A | 1 | +15 V pwr |
| B | 6, 3 | Pwr, signal gnd |
| C | 2 | -15 V pwr |
| D | 7 | Tilt out, ±5V = ±1 degree |

Converting tilts to leg errors

The height of the west leg is fixed.   Designate the height of the north and south legs as

N = W + ΔN, S = W + ΔS.   If we express the tilt as a function of azimuth as:

t(φ) = A cos(φ) + B sin(φ) + C

Where

A = (ΔN + ΔS) / 2d1 ,

B = (ΔS – ΔN) / d2

and C is a constant offset term.  One can derive A and B from a least squares fit to the tilt vs. azimuth, or more crudely, from the tilts at azimuths of 0, 90, 180, and 270 degrees:

A = 0.5 ( t(0) – t(180) )

B = 0.5 ( t(90) – t(270) ).

Thus, if the tilt is more positive at φ=90 than at φ=270, the south leg is higher than the north leg.  Solving for ΔN and ΔS in inches, one finds:

ΔN (inches) = 0.024 A (arcmin) – 0.014 B (arcmin)

ΔS (inches) = 0.024 A (arcmin) + 0.014 B (arcmin)

One must remove these errors in order to level the antenna; that is, if ΔN = 0.1 inch, one must shorten the N leg by 0.1 inch.  The adjustment screws have 4 threads per inch, so one full turn changes the height by 0.25 inch.

Adjusting the feet

- Loosen tiedown nuts at the 4 corners of the base.

- Jack up the N or S corner of the base to take the weight off the foot.  Be sure to insert the jack under the reinforcing plate – see diagram below.

- Use special wrench to turn adjustment ring on the foot.  Turning the ring CW unscrews the support screw from the base, thus raises the base.

- Lower the jack; adjust other leg; remeasure tilt.

- When finished, tighten the tiedowns at the 4 corners of the base.