Installing WBA13 bias boards into rcvr dewar control box

Note: there are 2 versions of the board 765C, the dewar control board.  One has green solder mask, the other does not.  Wiring is very similar, though not identical.

- Remove bottom lid of dewar control box.

- Drill 4 mounting holes in bottom lid for the bias board.

- Remove resistors: R21, R25, R36, R37, R17, R20, R33, R40.

- Remove ICs: U37, U29.

- Cut traces between R21 and U18, between R25 and U18.

- If test points are present on front of box, remove the left-hand test point so LED can be installed in the hole; wire to the test point can be removed or coiled up.

- If test points are not present on front of box, drill a .25" diam hole to accomodate the LED.

- Wire Molex connectors as specified in attached spreadsheet.

- Replace bottom lid, mount bias board in box, connect Molex connectors, install LED into hole in front panel.

- If there are other test points, solder wire from test point just to the right of the LED to TP2 (Vdrain mon), next test point to TP3 (Vgate mon).

- Reinstall dewar control box in telescope, connect all cables EXCEPT cable 204; install 37-pin D test plug instead.   Confirm that readouts through atod are correct, then power down, install cable 204.